## Main
Ticket No. | Ticket Type | Business Unit | Team Manager | Concerned Engineer | Warranty Status | Customer Name | Customer Address | Valid/Invalid | Customer Complains | Mobile No | Email | Status | Start Date | Due Date | Updated Date | Mode of Support | Priority | Action Taken | Latest Update | Reason For Delay | Closing Date | Brief Solution | MoM Document
230009 | Equipment Complaint | ATE | Keyur Makadia |  | Out of Warranty | Dynalog India | Kailash Complex, Park Site, Vikhroli (W), Mumbai 400 079. Name: - Ganesh Bankhele | Valid | Need to visit site for 1 day to know the concern | 7045709698 | ganeshb@dynalogindia.com | Closed | Jul 6, 2023 | Jul 13, 2023 | Jul 7, 2023 12:59 AM | Telephonic/ Email
230008 | Support | ATE | Keyur Makadia | Kartik Nattar | In Warranty | Soham | SAC Ahmedabad | Valid | facing problem in LabVIEW and OST card. | 07926914218 | soumen@sac.isro.gov.in | Closed | Jun 26, 2023 | Jul 3, 2023 | Aug 12, 2023 12:10 AM |  | Medium | User was using LV2020 for controlling OST module but the module is not supported for LV2020 or higher. Told user to use LV 2018. User first tried with LV2019 and OST module |  |  |  | OST module is currently supported in LV2019 or lower version
230006 | Support | ATE | Keyur Makadia | Application Engineer | In Warranty | Dheeraj Sharma | RRCAT Indore | Valid | This is regarding the issue we have discussed yesterday along with one of your NI colleagues. I am attaching a snapshot of NI Max indicating all the software components I have installed on a PXIe-8135 controller seating on the PXIe chassis 1082. This chassis also have NI 7975R FlexRIO and NI 5783 adapter module. All these hardwares are not showing up in the "Devices & Interfaces" tree. Kindly suggest. | 9009451141 | dheerajsharma@rrcat.gov.in | In Progess | May 16, 2023 | May 16, 2023 | May 28, 2023 11:20 PM |  | High
230005 | Equipment Complaint | Sales | Dhavalsinh Chauhan |  | Out of Warranty | CNH, Indore | Indore | Valid | Fluke DS703FC Battery replacement | 8826063118 | satyendra.thakur@external.cnhind.com | Open | Apr 25, 2023 | May 31, 2023 |  | Telephonic/ Email
230004 | Equipment Complaint | Sales | Dhavalsinh Chauhan |  | In Warranty | Newen systems pvt ltd | Makarpura Baroda | Valid | Fluke 80i-2010s Measurement issues | 7383909131 | pankush.anand@newensystems.com | Open | Apr 25, 2023 | May 17, 2023 |  | Telephonic/ Email
230003 | Equipment Complaint | Sales | Dhavalsinh Chauhan |  | In Warranty | Knack Energy PVT Ltd, Kadi | Kadi | Valid | Keithley 6517B Measurement Issues | 7984065842 | qc@knackenergy.com | Open | Apr 25, 2023 | May 16, 2023 |  | Telephonic/ Email
230002 | Equipment Complaint | Sales | Dhavalsinh Chauhan |  | In Warranty | Ammann India Private Limited | Mahesana | Valid | Fluke 930 Unit Not working | 9637998006 | Ketan.Bharambe@ammann.com | Open | Apr 25, 2023 | May 15, 2023 |  | Telephonic/ Email
220042 | Support | ATE | Keyur Makadia | Application Engineer | In Warranty | Manish K Dwivedi | SAC,ISRO, Ahmedabad. High Priority | Valid | Free LabVIEW Support | 07926913877 | mkdwivedi@sac.isro.gov.in | Closed | Dec 19, 2022 | Dec 19, 2022 | Jun 9, 2023 10:12 PM |  | High | The LABVIEW FULL DEVELOPMENT SYSTEM 2016 & LABVIEW FPGA MODULE 2016, have been installed & activated in one of the systems with the customer's serial number. The serial number can only be used to activate 2016 LabVIEW version. |  |  |  | We visited the lab on 1st March 2023 & activated the required LabVIEW License in PXI system
220040 | Support | ATE | Keyur Makadia | Ishan Patel | In Warranty | Space Application Centre | Ahmedabad | Valid | LabVIEW Support | 07926918234 | shrikantdubey@sac.isro.gov.in | Closed | Dec 12, 2022 | Dec 19, 2022 | Feb 10, 2023 4:40 PM |  | High | Today (12-12-2022) I will visit to resolve their issues faced in LabVIEW software development. | I visited on 12-12-22. Till now, I have resolved two queries. 1: They wanted to acquired LVDS data on external clock, store it in text file & measure the BER once acquisition is completed. I have made the FPGA & host VI to do the same. 2: They wanted to save the constellation data in CSV file. That is also completed Other two queries are open where we need help from Vinit so emailed him the queries. Waiting for reply |  |  | Customer issues were resolved by visiting their lab on 23-01-23 & remaining two open queries were resolved by NI.
220039 | Equipment Complaint | Sales | Dhavalsinh Chauhan | Dhavalsinh Chauhan | Out of Warranty | Sensewell Calibration Services | Vadodara | Valid | Temperature control issues.Fluke 4181 | 9725151582 | bhavin.shah@sensewellindia.com | Closed | Nov 22, 2022 | Mar 15, 2023 | Apr 21, 2023 10:57 PM | Site Visit | High | Customer has coordinated with Fluke service center | Customer has coordinated with Fluke service center and instrument is repair. |  | Mar 24, 2023
220036 | Equipment Complaint | Sales | Dhavalsinh Chauhan | Dhavalsinh Chauhan | Out of Warranty | Aarti Industries,Vapi | Vapi | Valid | Impedance measurement issues Fluke BT21 | 9727700692 | smit.naik@aarti-industries.com | Open | Nov 22, 2022 | Mar 15, 2023 | Nov 22, 2022 11:13 PM |  | High | Discussed with customer and unit has been sent to service centre | Discussed with customer and unit has been sent to service centre
220034 | Equipment Complaint | Sales | Dhavalsinh Chauhan | Dhavalsinh Chauhan | In Warranty | Prakash Deshmukh | BEL Bangalore | Valid | Measurement issue Fluke 287 | 9972655607 | prakashdeshmukh@bel.co.in | Closed | Sep 27, 2022 | Jan 31, 2023 | Nov 4, 2022 5:25 PM | Site Visit | High | Unit sent to service centre | Unit has benn repaired under warranty |  | Nov 2, 2022
220033 | Support | ATE | Keyur Makadia | Application Engineer | Out of Warranty | Jyotishankar mishra | IPR,Gandhinagar | Valid | A NI PXIe 1075 system 1. When powered the system runs automatically (push button not pressed) after 5-10sec. it turns off. Sometimes it automatically starts and shutdown. 2. Besides this also issues in OS ex-updates, python software, RAM. | 23964418 | jsmishra@ipr.res.in | Closed | Sep 19, 2022 | Sep 26, 2022 | Feb 10, 2023 4:17 PM |  | Medium | Following observation were stated: 1.For the system power on issue please refer this link & check the settings https://knowledge.ni.com/KnowledgeArticleDetails?id=kA00Z0000019WpNSAU&l=en-IN . If the settings are as per this link, then please do reverse settings & check if the problem is still there & let us know. 2.As per these links PXIe-8133 is obsolete & OS (Windows 7) is also obsolete https://www.ni.com/en-in/support/model.pxie-8133.html & https://www.ni.com/en-in/support/hardware-product-life-cycle-policies.html & https://www.ni.com/en-in/support/documentation/supplemental/19/national-instruments-support-for-windows-7--windows-server-2008-.html . Microsoft & NI has ended support for Windows 7, so upgradation & repairing is difficult. Also, RAM upgradation is not possible due to above reasons. Relevant RAM is also not available. 3.As per these links as PXIe-1075 is in maintenance phase https://www.ni.com/en-in/support/hardware-product-life-cycle-policies.html & https://www.ni.com/en-in/support/model.pxie-1075.html . So, there will be minimum support from NI. 4.Python is open-source software so please download from https://www.python.org/downloads/ as per your requirements. | Visited Customer & performed various troubleshooting step. Clean the modules & chassis. Reached to the conclusion that some modules are misbehaving when present in the chassis so as of now these modules are not inserted in the chassis. The problem seems to be resolved with this solution. Regarding RAM upgradation, it is not possible as relevant part no. is not available in NI. |  |  | Clean the modules & chassis. Some modules are not inserted in chassis & chassis is working fine.
220026 | Support | Sales | Dhavalsinh Chauhan | Dhavalsinh Chauhan | In Warranty | Inductotherm | Sanad Ahmedabad | Valid | Measurement Issues. 56180076 | 7228826836 | ndave@inductothermindia.com | Closed | Aug 18, 2022 | Nov 15, 2022 | Nov 4, 2022 7:56 PM | Site Visit | High | Unit is under diagnosis and therafter unit will be sent to service centre | Fluke I6000S FLEX-36 faulty probes have been replaced by new probe from service centre |  | Nov 4, 2022
220025 | Support | ATE | Keyur Makadia | Application Engineer | In Warranty | SAC ISRO | Ahmedabad | Valid | NI QyickSync support | 07926915348 | priya_mehrotra@sac.isro.gov.in | Closed | Aug 18, 2022 | Aug 31, 2022 | Feb 10, 2023 4:23 PM |  | Medium | Asked the customer to test the instrument with system having Windows 10 latest updates & all other required drivers & software related to QuickSyn Frequency Synthesizer. The customer was able to detect the instrument & configured according to the need. |  |  |  | The instrument was able to configured with system having Windows 10 latest updates & all other required drivers & software.
220020 | Support | Sales | Dhavalsinh Chauhan | Dhavalsinh Chauhan | In Warranty | Deepak Fertilizers Pvt Ltd | Deepak Fertilizer Pvt Ltd Dahej. | Valid | Discussed with user and Fluke 87 V is sent to service centre for repairing purpose | 9924018327 | instrument.dahej@dfpcl.com | Closed | Aug 9, 2022 | Oct 22, 2022 | Sep 1, 2022 2:23 PM | Site Visit | High | Discussed with user and Fluke 87 V is sent to service centre for repairing purpose | Unit has been repaired under warranty and sent to customer. |  | Aug 31, 2022
220017 | Support | ATE | Keyur Makadia | Application Engineer | In Warranty | Institute For Plasma Research | Gandhinagar | Valid | We purchased NI PXIe Data acquisition system consisting of 04 no. of NI PXIe-6378 modules (Ref. PO no. IPR/21-22/608074 dated 23/08/2021) We use these modules for Analog Input channels since last few weeks. Out of these 04 modules, one module with Sr. no. 198927D-05L 1F47540 is malfunctioning. The said module is not acquiring any of 16 AI channels properly. Although we are able to detect the module in NI-MAX utility. The module is in warranty period. Kindly do the needful to troubleshoot the issue. | 07923962020 | imran@ipr.res.in | Closed | Jul 5, 2022 | Jul 31, 2022 | Jul 25, 2022 7:01 PM |  | Medium | 1.) Checked the PXIe 6378 card with NI MAX Self Calibration test. The card failed the test. 2.) Then checked with LabVIEW program but the acquisition was not as per the requirement. 3.) Raised the issue in NI Technical Support Service. | NI Technical support team confirmed hardware failure & they raised the ticket with their RMA team for reparation. Card need to be taken out for further investigation. Repair request has been raised in NI. |  |  | After investigation it was found that module is not functioning as per need & need to be repair. Reparation request has been raised in NI.
220015 | Support | Sales | Dhavalsinh Chauhan | Dhavalsinh Chauhan | In Warranty | Grasim Industries Limited, Nagda | Nagda | Valid | Fluke-376- 45480253WS-Measurement issues | 9977448257 | satish.baghmar@adityabirla.com | Closed | Jun 29, 2022 | Sep 30, 2022 | Dec 5, 2022 6:51 PM | Site Visit | High | Discussed the issues with user | Fluke 376 Sr No: 45480253WS has been repaired under warranty. We are sending the unit on 5-12-22 |  | Dec 5, 2022
220005 | Support | Sales | Dhavalsinh Chauhan | Dhavalsinh Chauhan | In Warranty | Inductotherm | Inductotherm Ahmedabad Mitul Bhatt | Valid | Waveforms are observed noisy while connecting to load according to customer in Tektronix TBS1072B | 7228826836 | testheat@inductothermindia.com | Closed | May 11, 2022 | May 18, 2022 | Jun 9, 2022 6:45 PM | Telephonic/ Email | High | observed the waveforms and sent the oscilloscope to service centre | Oscilloscope was repaired under warranty and working ok |  | May 23, 2022