## Main
Ticket No. | Ticket Type | Business Unit | Team Manager | Concerned Engineer | Warranty Status | Customer Name | Customer Address | Valid/Invalid | Customer Complains | Mobile No | Email | Status | Start Date | Due Date | Updated Date | Mode of Support | Priority | Action Taken | Latest Update | Reason For Delay | Closing Date | Brief Solution | MoM Document
240006 | Support | ATE | Keyur Makadia | Kartik Nattar | In Warranty | Shrikant Dubey | SAC Bopal Capmus | Valid | Support in Migration of NI 7975R to NI 7932 R Code and installation of required drivers. | 07926918675 | shrikantdubey@sac.isro.gov.in | Closed | Jan 31, 2024 | Feb 10, 2024 | Feb 5, 2024 9:46 PM |  | Medium | Installed LabVIEW 2018 SP1 in end user PC. Change the FPGA target in existing project with 7932R and complied the code. Tested the code with loop back |  |  |  | Change the FPGA target in existing project with 7932R and complied the code.
240005 | Support | ATE | Keyur Makadia |  | In Warranty | asdf | asdf | Valid | asdf | 9876543210 | sdfg@gmail.com | Open | Jan 31, 2024 | Feb 15, 2024 |  | Telephonic/ Email
230009 | Equipment Complaint | ATE | Keyur Makadia |  | Out of Warranty | Dynalog India | Kailash Complex, Park Site, Vikhroli (W), Mumbai 400 079. Name: - Ganesh Bankhele | Valid | Need to visit site for 1 day to know the concern | 7045709698 | ganeshb@dynalogindia.com | Closed | Jul 6, 2023 | Jul 13, 2023 | Jul 7, 2023 12:59 AM | Telephonic/ Email
230008 | Support | ATE | Keyur Makadia | Kartik Nattar | In Warranty | Soham | SAC Ahmedabad | Valid | facing problem in LabVIEW and OST card. | 07926914218 | soumen@sac.isro.gov.in | Closed | Jun 26, 2023 | Jul 3, 2023 | Aug 12, 2023 12:10 AM |  | Medium | User was using LV2020 for controlling OST module but the module is not supported for LV2020 or higher. Told user to use LV 2018. User first tried with LV2019 and OST module |  |  |  | OST module is currently supported in LV2019 or lower version
230006 | Support | ATE | Keyur Makadia | Application Engineer | In Warranty | Dheeraj Sharma | RRCAT Indore | Valid | This is regarding the issue we have discussed yesterday along with one of your NI colleagues. I am attaching a snapshot of NI Max indicating all the software components I have installed on a PXIe-8135 controller seating on the PXIe chassis 1082. This chassis also have NI 7975R FlexRIO and NI 5783 adapter module. All these hardwares are not showing up in the "Devices & Interfaces" tree. Kindly suggest. | 9009451141 | dheerajsharma@rrcat.gov.in | In Progess | May 16, 2023 | May 16, 2023 | May 28, 2023 11:20 PM |  | High
220042 | Support | ATE | Keyur Makadia | Application Engineer | In Warranty | Manish K Dwivedi | SAC,ISRO, Ahmedabad. High Priority | Valid | Free LabVIEW Support | 07926913877 | mkdwivedi@sac.isro.gov.in | Closed | Dec 19, 2022 | Dec 19, 2022 | Jun 9, 2023 10:12 PM |  | High | The LABVIEW FULL DEVELOPMENT SYSTEM 2016 & LABVIEW FPGA MODULE 2016, have been installed & activated in one of the systems with the customer's serial number. The serial number can only be used to activate 2016 LabVIEW version. |  |  |  | We visited the lab on 1st March 2023 & activated the required LabVIEW License in PXI system
220040 | Support | ATE | Keyur Makadia | Ishan Patel | In Warranty | Space Application Centre | Ahmedabad | Valid | LabVIEW Support | 07926918234 | shrikantdubey@sac.isro.gov.in | Closed | Dec 12, 2022 | Dec 19, 2022 | Feb 10, 2023 4:40 PM |  | High | Today (12-12-2022) I will visit to resolve their issues faced in LabVIEW software development. | I visited on 12-12-22. Till now, I have resolved two queries. 1: They wanted to acquired LVDS data on external clock, store it in text file & measure the BER once acquisition is completed. I have made the FPGA & host VI to do the same. 2: They wanted to save the constellation data in CSV file. That is also completed Other two queries are open where we need help from Vinit so emailed him the queries. Waiting for reply |  |  | Customer issues were resolved by visiting their lab on 23-01-23 & remaining two open queries were resolved by NI.
220033 | Support | ATE | Keyur Makadia | Application Engineer | Out of Warranty | Jyotishankar mishra | IPR,Gandhinagar | Valid | A NI PXIe 1075 system 1. When powered the system runs automatically (push button not pressed) after 5-10sec. it turns off. Sometimes it automatically starts and shutdown. 2. Besides this also issues in OS ex-updates, python software, RAM. | 23964418 | jsmishra@ipr.res.in | Closed | Sep 19, 2022 | Sep 26, 2022 | Feb 10, 2023 4:17 PM |  | Medium | Following observation were stated: 1.For the system power on issue please refer this link & check the settings https://knowledge.ni.com/KnowledgeArticleDetails?id=kA00Z0000019WpNSAU&l=en-IN . If the settings are as per this link, then please do reverse settings & check if the problem is still there & let us know. 2.As per these links PXIe-8133 is obsolete & OS (Windows 7) is also obsolete https://www.ni.com/en-in/support/model.pxie-8133.html & https://www.ni.com/en-in/support/hardware-product-life-cycle-policies.html & https://www.ni.com/en-in/support/documentation/supplemental/19/national-instruments-support-for-windows-7--windows-server-2008-.html . Microsoft & NI has ended support for Windows 7, so upgradation & repairing is difficult. Also, RAM upgradation is not possible due to above reasons. Relevant RAM is also not available. 3.As per these links as PXIe-1075 is in maintenance phase https://www.ni.com/en-in/support/hardware-product-life-cycle-policies.html & https://www.ni.com/en-in/support/model.pxie-1075.html . So, there will be minimum support from NI. 4.Python is open-source software so please download from https://www.python.org/downloads/ as per your requirements. | Visited Customer & performed various troubleshooting step. Clean the modules & chassis. Reached to the conclusion that some modules are misbehaving when present in the chassis so as of now these modules are not inserted in the chassis. The problem seems to be resolved with this solution. Regarding RAM upgradation, it is not possible as relevant part no. is not available in NI. |  |  | Clean the modules & chassis. Some modules are not inserted in chassis & chassis is working fine.
220025 | Support | ATE | Keyur Makadia | Application Engineer | In Warranty | SAC ISRO | Ahmedabad | Valid | NI QyickSync support | 07926915348 | priya_mehrotra@sac.isro.gov.in | Closed | Aug 18, 2022 | Aug 31, 2022 | Feb 10, 2023 4:23 PM |  | Medium | Asked the customer to test the instrument with system having Windows 10 latest updates & all other required drivers & software related to QuickSyn Frequency Synthesizer. The customer was able to detect the instrument & configured according to the need. |  |  |  | The instrument was able to configured with system having Windows 10 latest updates & all other required drivers & software.
220017 | Support | ATE | Keyur Makadia | Application Engineer | In Warranty | Institute For Plasma Research | Gandhinagar | Valid | We purchased NI PXIe Data acquisition system consisting of 04 no. of NI PXIe-6378 modules (Ref. PO no. IPR/21-22/608074 dated 23/08/2021) We use these modules for Analog Input channels since last few weeks. Out of these 04 modules, one module with Sr. no. 198927D-05L 1F47540 is malfunctioning. The said module is not acquiring any of 16 AI channels properly. Although we are able to detect the module in NI-MAX utility. The module is in warranty period. Kindly do the needful to troubleshoot the issue. | 07923962020 | imran@ipr.res.in | Closed | Jul 5, 2022 | Jul 31, 2022 | Jul 25, 2022 7:01 PM |  | Medium | 1.) Checked the PXIe 6378 card with NI MAX Self Calibration test. The card failed the test. 2.) Then checked with LabVIEW program but the acquisition was not as per the requirement. 3.) Raised the issue in NI Technical Support Service. | NI Technical support team confirmed hardware failure & they raised the ticket with their RMA team for reparation. Card need to be taken out for further investigation. Repair request has been raised in NI. |  |  | After investigation it was found that module is not functioning as per need & need to be repair. Reparation request has been raised in NI.